Sample Engineering Specification

Talon HVX Field Pour, Above-Grade, Full Electronic Truck Scale

- PART 1 - GENERAL

  - Scope.  This section sets forth the requirements for one above-grade truck scale. The scale shall be furnished and installed complete as specified in the following paragraphs, including platform, foundation, load cells, instruments, surge voltage protection system, and the services of the manufacturer’s service representative.

    - Acceptable Manufacturer.  The truck scales furnished under this section shall be manufactured by Fairbanks Scales or equal.

  - General.  Equipment furnished and installed under this section shall be assembled, erected, and placed in proper operating condition in full conformity with drawings, specifications, engineering data, instructions, and recommendations of the equipment manufacturer unless exceptions are noted by the Engineer.

    - General Equipment Stipulations.  The General Equipment Stipulations shall apply to all equipment furnished under this section.

    - Governing Standards.  The scale systems shall have been issued a Certificate of Conformance by the National Type Evaluation Program, (N.T.E.P.) and shall conform to the following federal, state, local, and industrial standards.

      - National Institute of Standards and Technology (NIST), Handbook 44, "Specifications, Tolerances, and Technical Requirements for Weighing and Measuring Devices".

      - Applicable state regulations for commercial weighing devices.

      - American Welding Society AWS D1.1, 2006.

      - Weighbridge structural steel and structural steel embedments shall conform to ASTM A36.

      - Foundation reinforcing steel shall conform to ASTM A615, Grade 40 or 60.

    - Anchor Bolts.  All anchor bolts, nuts, and washers shall be made of carbon steel or comparable, and may be plated.

    - Edge Grinding.  Sharp projections of cut or sheared edges of ferrous metals, which are not to be welded, shall be ground as required to ensure paint adherence.

    - Surface Preparation.  All ferrous metal surfaces, except stainless steel or plated steel, shall be blast cleaned in the shop in accordance with the paint manufacturer’s recommendations.  All mill scale, rust, and contaminants shall be completely removed before shop primer is applied. The components of each module shall be cleaned to an SSPC-SP6 finish prior to painting.

    - Shop Painting.  All steel surfaces, except stainless steel or plated steel, shall be coated with a PPG AUE-370 high solids acrylic urethane, a two component chemically cross-linked commercial coating system.  The paint system will be applied per the manufacturer’s recommendations.  See attached data sheet.

Additional field painting other than touchup painting of damaged surfaces will not be required.

    - Power Supply.  Unless otherwise specified, the power supply to the equipment will be a dedicated 120 volt, single phase, 60 Hz connection.  Where control voltages lower than the power supply voltage is required, suitable control power transformers shall be furnished.

    - Surge Voltage Protection.  The scale system shall be surge and lightning protected.  This protection shall include a surge protection device which plugs into a standard 115 VAC outlet. The load cells shall be optically isolated, and surge protected. The scale shall have a dual point ground rod system for the grounding of the weighbridge, power supply, and the sectional controllers. Electrical diagrams of the scale grounding and surge protection shall be supplied with submittals.  The surge voltage package shall be provided as a unit and be tested and approved by the scale manufacturer.

Surge protection devices or components not designed or tested by the scale manufacturer as a unit are unacceptable.

  - Submittals.  Complete foundation and installation drawings, together with detailed specifications and data covering materials, parts, devices, and accessories forming a part of the equipment furnished, shall be submitted in accordance with the submittals section.  Drawings shall cover all scale components, foundation details, and pier loading information necessary for the design of the scale foundation or installation.

- PART 2 – PRODUCTS

  - Scale Design

    - General Description. The scale shall be a Fairbanks model Talon Series, above-grade, fully-electronic truck scale with rocker column load cells. Scale shall be a “flat top” orthotropic design and not require side beams above the surface of the deck. Equipment is to consist of parts designed to act as a unit by a manufacturer experienced in design, construction, manufacture of electronic components, as well as operation of equipment for the purpose required.

    - Scale Capacity and CLC. The N.T.E.P. approved scale shall have a minimum capacity of ____________ tons and have a minimum concentrated load capacity (CLC) of 100,000 lbs.

    - Scale eMin. The N.T.E.P. approved scale shall have an eMin (as defined by NIST HB 44) approval of 10 lbs.  The current NTEP Certificate of Conformance shall be provided as part of the project documentation package to support this approval.

    - Weighbridge Design. The platform shall be _______ feet long and _______ feet wide and be comprised of 12” concrete deck reinforced by 6” structural steel wide flange beam stringers which run the full length of each module.  Side beams shall be 12” structural steel members having nelson studs attached to the internal sides of each beam to create a composite connection between the concrete and the steel. Bottom plates of the module will be ¼” smooth plate to which the above mentioned structural steel stringers will be welded. Scale shall not require access plates on deck surface to service load cells or junction boxes. Removable cover plates are not acceptable on deck surface. Understructure components consisting of formed U’s or bent plates shall not be used. The bottom plate is to be welded to the flange support from the inside of the structure.

The maximum deflection for legal highway loads shall not exceed 1/1300th of span and the bending stress in any member shall not exceed 20,000 psi at the rated CLC of 100,000 lbs.

    - Checking. The scale shall utilize longitudinal and lateral bumper checking. The scale shall be completely self-checking. No check rods or flexure checking plates are to be used. Checking shall be mechanically isolated from the load cell and load cell stands/base plates.

    - Load Cell and Controller/Junction Box Specifications. The scale shall have rocker column load cells.  Each rocker column load cell shall have a minimum gross rated capacity of 110,000 lbs each.  The load cells shall be constructed of stainless steel and shall be completely hermetically sealed to prevent moisture from entering the cell. Load cells shall be nitrogen filled to retard internal corrosion. Load cells shall meet the following specifications:

Design	Rocker column

Load cell construction:

Element 	420 Stainless steel

Can	304 Stainless steel

Minimum capacity:	110,000 lbs

Maximum Height:	4.66”

Protection: 	NEMA 6P/IP69K

Excitation voltage: 	4 to 24 Volts

Input resistance (ohms): 	1080 +/- 10

Output resistance (ohms): 	1080 +/- 1

Insulation resistance (megohms): 	≥5000

Operating temperature: 	-40 to 203oF (-40 to 95oC)

Temperature effect on sensitivity (%/°C): 	≤ 0.0007%

Safe load limit (% of Capacity): 	200%

Ultimate load (% of Capacity): 	300%

Minimum dead load: 	0

Rated Output (mV/V): 	2.4

Load cell cable:	Stainless steel braid

Sealing: 	Hermetically sealed; gas filled, cable entry sealed by glass to metal header

NTEP Accuracy Class: 	IIIL, 10,000 divisions

The load cell shall have an available 25 year warranty.

The load cell shall be provided with a shielded, 4 conductor cable with a polyurethane jacket. Load cell cabling shall be protected by a stainless steel braided armor to protect against abrasion or rodent damage.  Load cell cables must be hard wired to the load cell.  “Quick disconnect” load cell cable connectors on cells or on junction boxes are not acceptable.

Load cell output shall be a DC analog signal. Load cells with digital output shall not be acceptable. Each load cell shall be optically isolated via a smart sectional controller (a junction board that converts the analog load cell signal into a digital signal via fiber optics or Intalogix communication). PC boards will be encapsulated in epoxy or plastic. A board that is not protected in this fashion is unacceptable. Furthermore, each encapsulated board shall be housed in a type 304 Stainless steel enclosure rated NEMA 4X.

The scale shall have self diagnostic capabilities able to identify load cell problems, failure, and predict failure before it occurs to prevent downtime. The diagnostics are to measure load cell counts (not weight) and will be used to determine reliability. Should a load cell fail, the instrumentation shall identify the specific load cell that has failed. All trouble shooting shall be done from within the scale house.

  - Platform and Foundation Requirements. The weighbridge and load cell assemblies shall be supported by a reinforced concrete pier type or full slab foundation as indicated on the drawings. The dimensions for the scale foundation and platform shall be as recommended by the equipment manufacturer and accepted by the Engineer.  Reinforcing steel placement and structural steel embedment placement shall be performed as shown on the manufacturer's foundation drawings.

The scale manufacturer shall furnish the following items for construction of the scale platforms and pits:

- Weighbridge modules

- Longitudinal and lateral bumper checking devices

- Load cells and load cell assemblies

- Anchor bolts

- Platform and endwall painted structural steel embedments

- PART 3 – EXECUTION

  - Installation.  The scale shall be manufactured, provided, and installed by a scale company that has a minimum of five years of experience installing similar truck scale systems.

The installer shall configure the scale system as indicated on the certified drawings.  All concrete work shall be as specified in the cast-in-place concrete section.  Anchor bolts shall be set as required by the scale manufacturer’s drawings.

  - Manufacturer’s Field Services.  Where scheduled in the equipment schedule section, an experienced, competent, and authorized representative of the manufacturer shall provide field services for equipment furnished under this section.  Field services shall meet the requirements of Manufacturer's Field Services in the quality control section.

  - Field Testing and Acceptance.  An authorized manufacturer’s representative shall provide the required scale certification for capacity and accuracy to the Engineer as required by the applicable state department of weights and measures and any other applicable state or county agency.

  - Personnel Training.  An experienced, competent, and authorized representative of the manufacturer shall train the Owner's personnel in operating, maintaining, and repairing the equipment specified in this section.  The training provided shall meet the requirements of Personnel Training Services in the quality control section.  The number of training sessions and duration of each session shall be as scheduled in the equipment schedule section.

End of Section