# Sample Engineering Specification

# Aegis® Stainless Steel Drum Scale

- PART 1 - GENERAL

  - Scope.  Furnish and install a fully electronic, stainless steel drum scale with a capacity of 1000 pounds, platform size of 30 inches by 30 inches, and three-inch wide integral access ramps for entrance from either end of the scale.

    - Acceptable Manufacturer.  Equipment furnished under this section shall be manufactured by Fairbanks Scales or equal.

    - Acceptable Installation Execution.  Equipment furnished under this section shall be installed by manufacturer’s factory-trained technicians.

  - General.  Equipment furnished and installed under this section shall be assembled, erected, and placed in proper operating condition in full conformity with drawings, specifications, engineering data, instructions, and recommendations of the equipment manufacturer unless exceptions are noted by the Engineer.

    - General Equipment Stipulations.  The General Equipment Stipulations shall apply to all equipment furnished under this section.

    - Governing Standards.  The scale system shall have been issued a Certificate of Conformance by the National Type Evaluation Program, (NTEP) and shall conform to the following federal, state, local, and industrial standards.

      - National Institute of Standards and Technology (NIST), Handbook 44, "Specifications, Tolerances, and Technical Requirements for Weighing and Measuring Devices".

      - Applicable state regulations for commercial weighing devices.

      - American Welding Society AWS D1.1, current edition.

      - Weighbridge structural stainless steel and structural stainless steel embedment shall conform to ASTM A480/A480M and A240.

    - Power Supply.  Unless otherwise specified, the power supply to the equipment will be a dedicated 120 Volt, single phase, 60 Hz connection.  Where control voltages lower than the power supply voltage is required, suitable control power transformers shall be furnished.

    - Surface Preparation.  All austenitic stainless steel shall be cleaned and the finish wire brushed per shop standards.

- PART 2 – PRODUCTS

  - Scale Design

    - General Description.  The scale shall be a Fairbanks model Aegis Stainless Steel Drum Scale or equal with digital indicator as specified in paragraph 2.1.6. The scale will be suitable for use on a reasonably flat (+/- 3˚) concrete floor. Equipment is to consist of parts designed to act as a unit by a manufacturer experienced in design, construction, manufacture of electronic components,        and operation of equipment for the purpose required.

    - Scale Capacity and Platform Size.  The NTEP approved scale shall have a capacity of 1,000 pounds and shall be 30 inches long and 30 inches wide.

    - Platform Structure.  The platform will be a self-contained design. The platform will be constructed of 304 stainless steel smooth plate. All structural steel portions of the scale, ramps and associated components will have a similar finish. The platform structure will consist of a four load cell fully electronic design and include lifting eyeholes to allow cleaning underneath. The corner suspension will be designed to keep the scale assembled when it is lifted by the platform. The platform weighing surface height will be no more than 1.5” inches.

    - Load Cells.  Single-ended, shear beam, strain gage load cells will be used to support each corner of the platform in a manner that prevents the possibility of platform tipping. Cells will have stainless steel construction and an IP69K rating.

    - Live Side Rails.  Live side rails enable the accurate weighing of bulk products and oversized containers, such as pallets and totes.

    - Digital Instrument.  The digital instrument will be Fairbanks model FB 2250 Series or equal and must include the following standard features as a minimum requirement:

      - Display with a minimum of 1.0” high characters.

      - Auto tare, single pushbutton for setting tare weight of an empty container into instrument memory.

      - Units selection button for weigh mode setup (lb, kg, g, oz, lb-oz, or custom).

      - Front panel legends to indicate weigh mode (net, lb, kg, oz, accumulation).

      - Auto zero tracking.

      - Stainless steel Housing (NEMA 4X rating minimum).

    - Accessories.  The specified Aegis Industrial Stainless Steel Drum Scale

includes the following accessories: Quantity of two (2), Aegis Stainless Steel

Drum Scale ramp, 30”.

End of Section