FAIRBANKS SCALES ENGINEERING SPECIFICATIONS FOR:

MODEL FB2550 NEMA 4X DIGITAL INSTRUMENT

## Section ____

### SCALE INSTRUMENTATION

PART 1 – GENERAL

- SCOPE.  This section covers the requirements for one digital scale instrument for use in the                               .  The scale instrument shall be furnished and installed complete as specified hereinafter, including the services of the manufacturer’s service representative.

    - Acceptable Manufacturer.  The scale instrumentation furnished under this section shall be manufactured by Fairbanks Scales or equal ISO accredited manufacturer.

- GENERAL.  Equipment furnished and installed under this section shall be assembled, calibrated and placed in proper operating condition in full conformity with wiring, specifications, engineering data, instructions and recommendations of the equipment manufacturer unless exceptions are noted by the Engineer.

    - General Equipment Stipulations.  The General Equipment Stipulations shall apply to all equipment furnished under this section.

    - Governing Standards.  The scale systems shall have been issued a Certificate of Compliance by the National Type Evaluation Program, (N.T.E.P.) and shall conform to the following federal, state, local, and industrial standards.

    - National Bureau of Standards, NBS Handbook 44, "Specifications, Tolerances, and Technical Requirements for Weighing and Measuring Devices".

    - National Electrical Manufacturers Association and the National Electrical Code.

    - Applicable state regulations for commercial weighing devices.

    - Power Supply.  Unless otherwise specified, the power supply to the equipment will be a dedicated 120 volts, single phase, 60 Hz.  Where control voltages lower than the power supply voltage is required, suitable control power transformers shall be furnished.

PART 2 – PRODUCTS

  - SCALE INSTRUMENT DESIGN.  The scale instrument shall be a Fairbanks Scales model FB2550 digital scale instrument or approved equivalent.

    - The scale instrument shall be N.T.E.P. approved with an Accuracy Class of III / IIIL.

    - The scale instrument shall feature a Category II Audit Trail.

  - SCALE INSTRUMENT HARDWARE SPECIFICATIONS

    - Processor shall be a 1.6 GHz processor minimum.

    - Memory shall be 1 GB RAM minimum.

    - Serial Ports shall be a minimum of three (3) externally accessible RS232 ports.

      - One RS232 port shall be capable of being configured for RS485.

    - Minimum of four (4) externally accessible USB 2.0 ports standard.

    - Dedicated 20 mA interface for interfacing with remote displays.  Interface shall support both active and passive remote displays.

    - External QWERTY keyboard interface shall support USB keyboards.

    - Ethernet Interface shall be through an external IEEE 802.3, 100Base-T compatible, 10/100 Mbps RJ45 connection.

    - Instrument shall be accessible through a standard PC browser using standard ethernet interface.  Use of any special software application for access instrument is prohibited.

      - Use of serial converters or similar means to obtain ethernet connectivity is prohibited.

    - The scale instrument display shall be a full color graphics display.

    - The scale instrument display shall provide minimum resolution of 800 x 480.

    - Display size shall be a minimum of 6.0” x 3.5” ( 7.0” diagonal)

    - Display shall be a TFT LCD with LED backlighting.

    - Display shall be an active matrix with touch screen standard.

    - The display shall be capable of displaying real time IP camera feed on the scale instrument weigh screen.  Up to two IP cameras can be interfaced to a single scale instrument.  Operator shall toggle between camera feeds by touching the on screen image.

    - On screen IP camera image shall be capable of being expanded to full instrument screen size by pressing a single key.

    - Still images can be captured and stored as part of the completed transaction record.

    - Still images shall be available for printing on transaction ticket.

    - The scale instrument shall display graphic image of vehicle on weigh screen when scale platform is loaded.  Image shall be full color graphic image, line drawing or outline drawing of a vehicle are not acceptable.  Vehicle type and color shall be selectable from a library of available images residing on the instrument.

    - The scale instrument shall include one external VGA Interface to provide connectivity to external VGA monitors.

    - The scale instrument load cell interfaces shall include support for Fairbanks Scales Intalogix Technology (digital interface for analog load cell communication) or standard analog load cell interfaces.   Use of proprietary load cells to obtain a digital interface is prohibited.

    - The scale instrument shall be able to detect water in scale pit if equipped with digital interface for load cell communication.  If detected, scale instrument shall display a message for the Operator on the weight screen.  Instrument shall also be capable of emailing notice of this condition to multiple recipients.

    - The scale instrument shall be capable of driving up to 36 load cells.

    - The scale instrument shall have the ability to control up to eight (8) individual scale platforms.  All eight platforms shall be visible on the weigh screen of the instrument.  The Operator shall switch between platforms by touching desired scale using the instrument touch screen.

    - The scale instrument shall have the ability to network up to five instruments together, using a standard ethernet network, each instrument controlling separate scale platform(s), automatically sharing transaction data, allowing any instrument to complete an open transaction started on a different instrument on the network.

    - The scale instrument shall be capable of interfacing to serial or USB printers for the purpose of producing tickets or reports.

    - Instrument shall be manufactured in the United States.

  - SCALE INSTRUMENT FUNCTIONS.  The scale instrument shall be a self-contained, operating unit providing a digital weight instrument driven by appropriate conditioning and control circuitry.  The scale instrument shall provide the capability to directly print a certified weight ticket.

    - The scale instrument shall feature a programmable Display rate with settings from 0.1 to 10 seconds in 0.1 second intervals.

    - The scale instrument shall feature programmable Zero button settings from 0 to 100% Zero capability.  Zero button shall also have the capability to be completely disabled.  Zero settings must also allow for a programmable Zero band threshold to be established based on weight, to facilitate a single print per weighment.

    - The scale instrument shall include a blind counter to track scale activity when platform is loaded, but no transaction is processed.

    - The instrument shall have the ability to automatically control traffic lights based on threshold weight settings, selectable time delays and transaction status.

    - The scale instrument shall have the ability to control combination remote display and traffic light devices through a single 20 mA interface.

    - The instrument shall have the ability to manually control traffic lights by touching the red/green traffic light icon on the instrument touch screen display.

    - The scale instrument shall have the ability to automatically assign an ID to a new Inbound transaction.

    - The scale instrument shall have the ability to store 500,000 individual tare weights.  Each stored Tare will include the Tare ID, unit of measure, date stored and vehicle description.

    - The scale instrument shall have the ability to store 250,000 individual Product or Material records.  Each stored Product will include the Product ID, unit of measure, net weight running total, Product description and two programmable conversion factors.

    - The scale instrument shall have the ability to display the stored Product list during the transaction by pressing a single key.

    - The scale instrument shall display the description of the selected Product on screen when processing a transaction.

    - The scale instrument shall have the ability to create Product Groups to associate and limit specific Products with specific Customer records.

    - The scale instrument shall have the ability to store 250,000 individual Customer records.  Each stored Customer will include the Customer ID, unit of measure, net weight running total, four lines of text for description, address, etc.

    - The scale instrument shall have the ability to display the stored Customer list during the transaction by pressing a single key.

    - The scale instrument shall display the four lines of description for the selected Customer on screen when processing a transaction.

    - The scale instrument shall have the ability to store 1,000,000 (one million) completed transactions or a minimum of 10,000 with two (2) images per transaction.

    - The scale instrument shall have the ability to display the Open/Inbound transaction list during the Outbound transaction by pressing a single key.

    - The scale instrument shall have the ability to create ten (10) additional transaction based Operator prompts for capturing additional information during the transaction process.  Data captured shall be available for printing on tickets and included in exported transaction data.

    - Transaction data can be printed or exported.

    - The scale instrument shall produce the following reports:

      - Completed Transactions by date range

      - Incomplete transactions by date range

      - Transactions sorted by Product ID, by date range

      - Transactions sorted by Customer ID, by date range

      - Summary by Product

      - Summary by Customer

      - Voided transactions

      - Scale activity

      - List of stored Tares

      - List of stored Products

      - List of Product groups

      - List of stored Customers

      - Calibration report

    - The scale instrument shall have the ability to export the following data:

      - Completed Transactions by date range

      - Incomplete transactions by date range

      - Transactions sorted by Product ID, by date range

      - Transactions sorted by Customer ID, by date range

      - Summary by Product

      - Summary by Customer

      - Voided transactions

      - Scale activity

      - List of stored Tares

      - List of stored Products

      - List of Product groups

      - List of stored Customers

      - Calibration report

    - Data shall be exportable in the following formats:

      - PDF

      - HTML

      - CSV

    - The scale instrument shall provide the means to turn data headers turned on and off in exported CVS data.

    - Data shall be accessible through a standard PC browser.

    - The instrument shall support the means to email exported transaction data.

    - The scale instrument shall support the means to automatically email a completed ticket once a transaction is completed.

    - The scale instrument shall have the ability to manually email a specific transaction, by ticket number.

    - The scale instrument shall have the means to capture up to two (2) static images from IP cameras.  Images are to be stored by ticket number.

    - Stored images shall be viewable by ticket number or by date range.  Individual images can be emailed or copied to a USB jump drive.

    - The scale instrument shall have the means to self diagnose and email the following conditions at a minimum:

      - Calibration Change

      - Float Switch Condition = ON

      - Load Cell Ghosted

      - Load Cell Failure

      - Routine Maintenance Required

      - Remote Access Enabled

      - Time and Date change

      - Scale Behind Zero

      - Blind Counter Increment

      - Instrument Configuration Change

      - Unauthorized Access Attempted

      - Sectional Error

      - Cell Motion Error

      - Zero Shift

      - Scale Capacity Exceeded

    - The scale instrument shall have the ability to drive eight (8) 1000 load cells at a distance of 1,800 feet.

    - The scale instrument shall visually ‘flag’ identified problem load cell(s) in diagnostic screen until flag is manually cleared to identify intermittent problems.

    - The scale instrument shall provide the means to view all load cells for a single platform at the same time to allow for proper scale system analysis. Individual viewing of cells is not acceptable.

    - The scale instrument shall be NTEP approved for wireless load cell communication.

    - The scale system shall have the ability to be 100% calibrated from the instrument.  No calibration adjustments or trimming shall be allowed at the scale platform.

    - The scale instrument shall have the means to calibrate the scale using the mV output of the connected load cells.

    - The scale instrument shall have the ability to electronically ‘ghost’ a failed load cell when equipped with the digital interface for load cell communication.

    - The scale instrument shall support remote troubleshooting by means of a standard PC browser. Use of any special software application to access the instrument is prohibited.

    - The scale instrument shall have the means to provide continuous weight data.  Weight data shall be provided serially or through the ethernet port.

    - The scale instrument shall feature the ability to backup to internal flash memory card or USB jump drive.

    - The scale instrument shall feature the ability to be backed up remotely using a standard PC browser.  Use of any special software application to backup instrument is prohibited.

PART 3 – EXECUTION

  - INSTALLATION.  The scale instrument shall be installed by a scale company that has a minimum of five years of experience installing similar scale systems.

The scale company must have an established service center in ____________ and have a current license on file with the local Weights & Measures Authority.

3-2.  MANUFACTURER'S FIELD SERVICES.  Where scheduled in the equipment schedule section, an experienced, competent, and authorized representative of the manufacturer shall provide field services for equipment furnished under this section.  Field services shall meet the requirements of Manufacturer's Field Services in the quality control section.

3-3.  FIELD TESTING AND ACCEPTANCE.  The authorized representative of the manufacturer shall provide the required scale certification for capacity and accuracy to the Engineer as required by the local Weights and Measures Authority any other applicable State or County agency.

3-4.  PERSONNEL TRAINING.  An experienced, competent, and authorized representative of the manufacturer shall train the Owner's personnel in operating and maintaining the equipment specified in this section.  The training provided shall meet the requirements specified in the quality control section.  The number of training sessions and duration of each session shall be as scheduled in the equipment schedule section.

End of Section