Sample Engineering Specification

AAR Deep Pit Style, Full Electronic Railroad Track/Truck Combination Scales

PART 1 - GENERAL

  - Scope.  This section sets forth the requirements for one full-electronic deep pit-style railroad track/motor truck combination scale. The scale shall be furnished and installed complete as specified in the following paragraphs, including weighbridge, load cells, instrumentation, platform, reinforcing steel and channels, along with pit coping, foundation bolts, nuts, and bumper plates should be furnished by the manufacturer.

    - Acceptable Manufacturer.  The truck/track scale(s) furnished under this section shall be manufactured by Fairbanks Scales or equal.

  - General.  Equipment furnished and installed under this section shall be assembled, erected, and placed in proper operating condition in full conformity with drawings, specifications, engineering data, instructions, and recommendations of the equipment manufacturer unless exceptions are noted by the Engineer.

    - General Equipment Stipulations.  The General Equipment Stipulations shall apply to all equipment furnished under this section.

    - Governing Standards.  The scale systems shall have been issued a Certificate of Conformance by the National Type Evaluation Program, (N.T.E.P.) and shall conform to the following standards:

      - National Institute of Standards and Technology (NIST), Handbook 44, "Specifications, Tolerances, and Technical Requirements for Weighing and Measuring Devices".

      - AREA/AAR requirements for railroad track scales

      - AREA/AAR Cooper E-80 bridge specifications

      - Applicable state regulations for commercial weighing devices.

      - American Welding Society AWS D1.1, latest edition.

      - Weighbridge structural steel and structural steel embedments shall conform to ASTM A36.

      - Reinforcing steel shall conform to ASTM A615, Grade 40 or Grade 60.

    - Anchor Bolts.  All anchor bolts, nuts, and washers shall be made of carbon steel or comparable, and shall be plated.

    - Edge Grinding.  Sharp projections of cut or sheared edges of ferrous metals, which are not to be welded, shall be ground as required to ensure paint adherence.

    - Surface Preparation.  All ferrous metal surfaces, except stainless steel or plated steel, shall be blast cleaned in the shop in accordance with the paint manufacturer’s recommendations.  All mill scale, rust, and contaminants shall be completely removed before shop primer is applied. The components of each module shall be cleaned to an SSPC-SP6 finish prior to painting.

    - Shop Painting.  All steel surfaces, except stainless steel or plated steel, shall be 	primed with a PPG ASP-795 Red Oxide Primer.  The primer will be applied per 	the manufacturer’s recommendations.  See attached data sheet.

Additional field painting other than touchup painting of damaged surfaces will not be required.

    - Power Supply.  Unless otherwise specified, the power supply to the equipment will be a dedicated 120 volt, single phase, 60 Hz connection.  Where control voltages lower than the power supply voltage are required, suitable control power transformers shall be furnished.

    - Surge Voltage Protection.  The scale system shall be surge and lightning protected.  This protection shall include a surge protection device which plugs into a standard 115 VAC outlet. The load cells shall be optically isolated, and surge protected. The scale shall have a dual point ground rod system for the grounding of the weighbridge, power supply, and the sectional controllers. The surge voltage package shall be provided as a unit and be tested and approved by the scale manufacturer.

  - Submittals.  Complete foundation and installation drawings, together with detailed specifications and data covering materials, parts, devices, and accessories forming a part of the equipment furnished, shall be submitted in accordance with the submittals section.  Drawings shall cover all scale components, foundation details, and pier loading information necessary for the design of the scale foundation or installation.

PART 2 – PRODUCTS

  - Scale Design

    - General Description. The scale shall be a Fairbanks Series 12-149X or equal. Scale platform assembly shall consist of reinforced concrete deck, steel weighbridge supported on multiple load cells.  Equipment is to consist of parts designed to act as a unit by a manufacturer experienced in design, construction, manufacture of electronic components, and operation of equipment for the purpose required.

    - Scale Capacity. The scale shall have a minimum scale capacity of 360 tons.

    - Sectional Capacity. The scale shall have a minimum sectional capacity of 180 tons.

    - Scale CLC. The N.T.E.P. approved scale shall have a minimum concentrated load capacity (CLC) of 80,000 lbs.

    - Weighbridge Design. The platform shall be _______ feet long and _______ feet wide and shall have a 6” reinforced concrete deck poured on-site using 4000 psi compressive strength concrete. The deck is to be lined along the bottom with corrugated steel, a reinforcing mat shall be set into place the length and width of the scale deck and the deck channel is to have studs welded to the steel to form a composite structure when the concrete is added.

    - Checking. The scale shall utilize rigid check rods to minimize deck movement.  Checking shall be mechanically isolated from the load cell, or load cell stands or base plates.

    - Load Cell and Controller/Junction Box Specifications. The load cell shall have a minimum capacity of 200,000 lbs. The load cell shall be provided with a shielded cable.

Minimum capacity:	200,000 lbs

Design	Compression Canister

Rated Output (mV/V): 	2 ± 0.1%

NTEP Accuracy Class: 	IIIL

Combined error (%FS): 	≤ ± 0.02500%

Temperature effect on sensitivity (%/°C): 	≤ ± 0.0072%

Excitation voltage: 	10 Nom. 25 Max.

Zero balance (%FS): 	≤ ± 5.0%

Compensated temperature range (°C):  	–10o to 40o

Operating temperature range (°C): 	–50o to 90o

Safe load limit (% of Capacity): 	150%

Ultimate load (% of Capacity): 	300%

Load cell material: 	Stainless load button and base

Sealing: 	Hermetically sealed

The scale shall have self-diagnostic capabilities able to identify load cell problems, failure, and predict failure before it occurs to prevent downtime. The diagnostics are to measure load cell counts (not weight) and will be used to determine reliability. Should a load cell fail, the instrumentation shall identify the specific load cell that has failed. All troubleshooting shall be done from within the scale house.

Load cell output shall be a DC analog signal. Load cells with digital output shall not be acceptable. Each load cell shall be optically isolated via a smart sectional controller. PC boards shall be encapsulated in epoxy or similar material. A board that is not protected in this fashion is unacceptable. Furthermore, each encapsulated board shall be housed in a type 304 stainless steel enclosure rated NEMA 4X. Access to the encapsulated board within the smart sectional controller enclosure shall be achieved without the use of tools.  Bolts, screws or other hardware shall not be used to seal the smart sectional controller enclosure.

  - Platform and Foundation Requirements. The weighbridge and load cell assemblies shall be supported by a reinforced concrete pier type or full slab foundation as indicated on the drawings. The dimensions for the scale foundation and platform shall be as recommended by the equipment manufacturer and accepted by the Engineer.  Reinforcing steel placement and structural steel embedment placement shall be performed as shown on the manufacturer's foundation drawings.

The scale manufacturer shall furnish the following items for construction of the scale platforms and pits:

- Weighbridge modules

- Longitudinal and lateral rigid checking devices

- Load cells and load cell assemblies.

- Anchor bolts.

- Painted structural steel embedments for platform and endwall coping.

- PART 3 – EXECUTION

  - Installation.  The scale shall be manufactured, provided, and installed by a scale company that has a minimum of five years of experience installing similar track scale systems.

The installer shall configure the scale system as indicated on the certified drawings.  All concrete work shall be as specified in the cast-in-place concrete section.  Anchor bolts shall be set as required by the scale manufacturer’s drawings.

  - Manufacturer’s Field Services.  Where scheduled in the equipment schedule section, an experienced, competent, and authorized representative of the manufacturer shall provide field services for equipment furnished under this section.  Field services shall meet the requirements of Manufacturer's Field Services in the quality control section.

  - Field Testing and Acceptance.  An authorized manufacturer’s representative shall provide the required scale certification for capacity and accuracy to the Engineer as required by the applicable state department of weights and measures and any other applicable state or county agency.

  - Personnel Training.  An experienced, competent, and authorized representative of the manufacturer shall train the Owner's personnel in operating, maintaining, and repairing the equipment specified in this section.  The training provided shall meet the requirements of Personnel Training Services in the quality control section.  The number of training sessions and duration of each session shall be as scheduled in the equipment schedule section.

End of Section