Sample Engineering Specification

Type S Cast Iron Mechanical Pit-Type Truck Scales

- PART 1 - GENERAL

  - Scope.  This section sets forth the requirements for one pit-type mechanical truck scale. The scale shall be furnished and installed complete as specified in the following paragraphs, including weighbridge, foundation, lever system, platform reinforcing, pit coping, load cells, indicators, surge voltage protection system, and the services of the manufacturer’s service representative.

    - Acceptable Manufacturer.  The truck scales furnished under this section shall be manufactured by Fairbanks Scales or equal.

  - General.  Equipment furnished and installed under this section shall be assembled, erected, and placed in proper operating condition in full conformity with drawings, specifications, engineering data, instructions, and recommendations of the equipment manufacturer unless exceptions are noted by the Engineer.

    - General Equipment Stipulations.  The General Equipment Stipulations shall apply to all equipment furnished under this section.

    - Governing Standards.  The scale systems shall have been issued a Certificate of Conformance by the National Type Evaluation Program, (N.T.E.P.) and shall conform to the following federal, state, local, and industrial standards.

      - National Institute of Standards and Technology (NIST), Handbook 44, "Specifications, Tolerances, and Technical Requirements for Weighing and Measuring Devices".

      - Applicable state regulations for commercial weighing devices.

      - American Welding Society AWS D1.1, 2006.

      - Weighbridge structural steel and structural steel embedments shall conform to ASTM A36.

      - Reinforcing steel shall conform to ASTM A615, Grade 40 or Grade 60.

    - Anchor Bolts.  All anchor bolts, nuts, and washers shall be made of carbon steel or comparable, and may be plated.

    - Edge Grinding.  Sharp projections of cut or sheared edges of ferrous metals, which are not to be welded, shall be ground as required to ensure paint adherence.

    - Surface Preparation.  All ferrous metal surfaces, except stainless steel or plated steel, shall be blast cleaned in the shop in accordance with the paint manufacturer’s recommendations.  All mill scale, rust, and contaminants shall be completely removed before shop primer is applied. The components of each module shall be cleaned to an SSPC-SP6 finish prior to painting.

    - Shop Painting.  All steel surfaces, except stainless steel or plated steel, shall be coated with a PPG AUE-370 high solids acrylic urethane, a two component chemically cross-linked commercial coating system.  The paint system will be applied per the manufacturer’s recommendations.  See attached data sheet.

All mechanical levers and suspension, shall be primed with a PPG ASP-795 Red Oxide Primer.  The primer will be applied per the manufacturer’s recommendations.  See attached data sheet.

Additional field painting other than touchup painting of damaged surfaces will not be required.

    - Power Supply.  Unless otherwise specified, the power supply to the equipment will be a dedicated 120 volt, single phase, 60 Hz connection.  Where control voltages lower than the power supply voltage is required, suitable control power transformers shall be furnished.

    - Surge Voltage Protection.  The scale system shall be surge and lightning protected.  This protection shall include a surge protection device which plugs into a standard 115 VAC outlet. The load cells shall be optically isolated, and surge protected. The scale shall have a dual point ground rod system for the grounding of the weighbridge, power supply, and the sectional controllers. Electrical diagrams of the scale grounding and surge protection shall be supplied with submittals.  The surge voltage package shall be provided as a unit and be tested and approved by the scale manufacturer.

Surge protection devices or components not designed or tested by the scale manufacturer as a unit are unacceptable.

  - Submittals.  Complete foundation and installation drawings, together with detailed specifications and data covering materials, parts, devices, and accessories forming a part of the equipment furnished, shall be submitted in accordance with the submittals section.  Drawings shall cover all scale components, foundation details, and pier loading information necessary for the design of the scale foundation or installation.

- PART 2 – PRODUCTS

  - Scale Design

    - General Description. The scale shall be a Fairbanks model Type “S” or equal. Scale platform assembly shall consists of reinforced concrete deck, steel weighbridge, and a cast iron lever system which delivers weight force to a single tension load cell. Equipment is to consist of parts designed to act as a unit by a manufacturer experienced in design, construction, manufacture of electronic components, and operation of equipment for the purpose required.

    - Scale Capacity and CLC. The N.T.E.P. approved scale shall have a minimum concentrated load capacity (CLC) of __________ lbs.

    - Weighbridge Design. The platform shall be _______ feet long and _______feet wide and shall have a 6” reinforced concrete deck poured on-site using 4000 psi compressive strength concrete. The deck is to be lined along the bottom with corrugated steel, a reinforcing mat shall be set into place the length and width of the scale deck, and the deck channel is to have studs welded to the steel to form a composite structure when the concrete is added. The scale shall have two 24” manholes for scales up to 45’, and three 24” manholes for scales 50’ and greater.

    - Lever System. The lever system shall be of rust-resistant cast iron construction with double web design. It shall utilize parallel link suspension to transmit weight from the weighbridge main girders to the cast iron levers and shall be free-floating and self-aligning, absorb the effects of platform motion on pivots and bearings, and reduce the effect of inertial shock caused by traffic.

There shall be adjustable vertical connections between all levers in the system. In addition, there should be elevation adjustment screws (two at each main lever) for vertical alignment of the scale deck with the concrete walls. There shall be removable bearings on each lever stand to provide for easy replacement of bearings without stand removal.

Pivots and bearings directly supporting the weight of the platform and the load applied to that platform shall be parallel with the direction of traffic. Pivots are to be supported over their entire length by self-aligning bearings, which shall be of high carbon steel and properly heat treated to a minimum Rockwell hardness of 58-62.

Cast iron lever system shall deliver reduced weight force to a single tension load cell. Load cell, junction boxes, and scale checking must be accessible and held in position by easily accessible hardware.

    - Checking. The scale shall utilize longitudinal and lateral bumper checking. The scale shall be completely self-checking. No check rods or flexure checking plates are to be used. Checking shall be mechanically isolated from the load cell, or load cell stands/base plates.

    - Load Cell and Controller/Junction Box Specifications. The scale shall have a load cell of an S-beam tension design.  The load cell shall be provided with a shielded cable.  If the scale is provided with Intalogix technology or similar, PC boards will be encapsulated in epoxy or similar material. A board that is not protected in this fashion is unacceptable. Furthermore, each encapsulated board shall be housed in a type 304 Stainless steel enclosure rated NEMA 4X. Access to the encapsulated board within the smart sectional controller enclosure shall be achieved without the use of tools.  Bolts, screws or other hardware shall not be used to seal the smart sectional controller enclosure.

  - Platform and Foundation Requirements. The weighbridge and load cell assemblies shall be supported by a reinforced concrete pier type or full slab foundation as indicated on the drawings. The dimensions for the scale foundation and platform shall be as recommended by the equipment manufacturer and accepted by the Engineer.  Reinforcing steel placement and structural steel embedment placement shall be performed as shown on the manufacturer's foundation drawings.

The scale manufacturer shall furnish the following items for construction of the scale platforms and pits:

- Weighbridge

- All components of lever system

- Longitudinal and lateral bumper checking devices

- Load cells and load cell assemblies

- Anchor bolts

- Platform and endwall painted structural steel embedments

- PART 3 – EXECUTION

  - Installation.  The scale shall be manufactured, provided, and installed by a scale company that has a minimum of five years of experience installing similar truck scale systems.

The installer shall configure the scale system as indicated on the certified drawings.  All concrete work shall be as specified in the cast-in-place concrete section.  Anchor bolts shall be set as required by the scale manufacturer’s drawings.

  - Manufacturer’s Field Services.  Where scheduled in the equipment schedule section, an experienced, competent, and authorized representative of the manufacturer shall provide field services for equipment furnished under this section.  Field services shall meet the requirements of Manufacturer's Field Services in the quality control section.

  - Field Testing and Acceptance.  An authorized manufacturer’s representative shall provide the required scale certification for capacity and accuracy to the Engineer as required by the applicable state department of weights and measures and any other applicable state or county agency.

  - Personnel Training.  An experienced, competent, and authorized representative of the manufacturer shall train the Owner's personnel in operating, maintaining, and repairing the equipment specified in this section.  The training provided shall meet the requirements of Personnel Training Services in the quality control section.  The number of training sessions and duration of each session shall be as scheduled in the equipment schedule section.

End of Section