Sample Engineering Specification

2000 Series, Full Electronic, Low Profile Railroad Track Scales

- PART 1 - GENERAL

  - Scope.  This section sets forth the requirements for one above-grade railroad track scale consisting of _____ weigh module(s). The scale shall be furnished and installed by the manufacturer complete as specified in the following paragraphs, including weighbridge, load cells, and instrumentation, along with foundation bolts, nuts, and bumper checking. All necessary rail hardware (excluding rail), anti-creep devices, and approach/dead span rail plates and anchor bolts shall be quoted separately, by the scale manufacturer, as optional equipment. Scale, dead span, and approach rails are not to be included and will be supplied by others (rail contractor or customer).

    - Acceptable Manufacturer.  The track scales furnished under this section shall be manufactured by Fairbanks Scales or equal.

  - General.  Equipment furnished and installed under this section shall be assembled, erected, and placed in proper operating condition in full conformity with drawings, specifications, engineering data, instructions, and recommendations of the equipment manufacturer unless exceptions are noted by the Engineer.

    - General Equipment Stipulations.  The General Equipment Stipulations shall apply to all equipment furnished under this section.

    - Governing Standards.  The scale systems shall have been issued a Certificate of Conformance by the National Type Evaluation Program, (N.T.E.P.) and shall conform to the following standards:

      - National Institute of Standards and Technology (NIST), Handbook 44, "Specifications, Tolerances, and Technical Requirements for Weighing and Measuring Devices".

      - AREA/AAR requirements for railroad track scales

      - AREA/AAR Cooper E-80 bridge specifications

      - Applicable state regulations for commercial weighing devices.

      - American Welding Society AWS D1.1, 2006.

      - Weighbridge structural steel and structural steel embedments shall conform to ASTM A36.

      - Reinforcing steel shall conform to ASTM A615, Grade 60.

    - Anchor Bolts.  All anchor bolts, nuts, and washers shall be made of carbon steel or comparable, and shall be plated.

    - Edge Grinding.  Sharp projections of cut or sheared edges of ferrous metals, which are not to be welded, shall be ground as required to ensure paint adherence.

    - Surface Preparation.  All ferrous metal surfaces, except stainless steel or plated steel, shall be blast cleaned in the shop in accordance with the paint manufacturer’s recommendations.  All mill scale, rust, and contaminants shall be completely removed before shop primer is applied. The components of each module shall be cleaned to an SSPC-SP6 finish prior to painting.

    - Shop Painting.  All steel surfaces, except stainless steel or plated steel, shall be primed with a PPG AUE-370 high solids acrylic urethane, a two component chemically cross-linked commercial coating system. The paint system will be applied per the manufacturer’s recommendations.  See attached data sheets.

Additional field painting other than touchup painting of damaged surfaces will not be required.

    - Power Supply.  Unless otherwise specified, the power supply to the equipment will be a dedicated 120 volt, single phase, 60 Hz connection.  Where control voltages lower than the power supply voltage is required, suitable control power transformers shall be furnished.

    - Surge Voltage Protection.  The scale system shall be surge and lightning protected.  This protection shall include a surge protection device which plugs into a standard 115 vac outlet. The load cells shall be optically isolated, and surge protected. The scale shall have a dual point ground rod system for the grounding of the weighbridge, power supply, and the sectional controllers. Electrical diagrams of the scale grounding and surge protection shall be supplied with submittals.  The surge voltage package shall be provided as a unit and be tested and approved by the scale manufacturer.

Surge protection devices or components not designed or tested by the scale manufacturer as a unit are unacceptable.

  - Submittals.  Complete foundation and installation drawings, together with detailed specifications and data covering materials, parts, devices, and accessories forming a part of the equipment furnished, shall be submitted in accordance with the submittals section.  Drawings shall cover all scale components, foundation details, and pier loading information necessary for the design of the scale foundation or installation.

- PART 2 – PRODUCTS

  - Scale Design

    - General Description. The scale shall be a Fairbanks 2000 Series Railroad Track Scale or equal. Scale will consist of either one module or two module placed end to end with a dead span between them to achieve the proper overall scale length. Equipment is to consist of parts designed to act as a unit by a manufacturer experienced in design, construction, manufacture of electronic components, and operation of equipment for the purpose required.

    - Scale Capacity and CLC. The N.T.E.P. approved scale shall have a sectional capacity of 170,000 lbs. Each 12’-6” module shall have a gross capacity of 170,000 lbs. Each 25’-0” module shall have a capacity of 340,000 lbs. Total gross capacity of any two modules used in combination shall be 340,000 lbs.

    - Weighbridge Design. The weighbridge assemblies shall be of factory-welded I-beam design. The platform shall have _______ sections, a total weigh length of _______ feet, and an overall scale rail length of _______feet. The dead span shall be _______ feet in length. The scale shall have a 1/4” thick, non-traffic bearing steel tread plate deck capable of supporting human traffic only, and no other vehicles other than rail cars will be driven over the scale.

It is imperative that vertical centerline of the scale rails, the vertical centerline of the main girders along their length, and the vertical and axial centerline of the load cells must all be aligned along the same vertical axis.

Each of the 12’-6” modules shall consist of a single prefabricated factory-welded weighbridge assembly with a non-traffic bearing steel tread plate. The 25’-0” module shall consist of two 12’-6” prefabricated factory-welded assemblies.

The scale must be designed to meet the requirements of A.R.E.A./A.A.R. specifications for maximum bending stress and maximum deflection under minimum Cooper E-80 loading characteristics.

    - Checking. The scale shall completely self-checking. No check rods shall be used. Checking shall be mechanically isolated from the load cell and load cell stands or base plates.

    - Load Cell and Controller/Junction Box Specifications. The load cells shall be constructed of stainless steel and shall be completely hermetically sealed to prevent moisture from entering the cell. Load cells shall be purged of air and then charged with dry nitrogen gas to retard internal corrosion. Load cells shall meet the following specifications:

Minimum capacity:	110,000 lbs

Design	Rocker column

Rated Output (mV/V): 	2 ± 0.1%

NTEP Accuracy Class: 	IIIL, 10,000 divisions

Combined error (%RO): 	≤ ± 0.0200%

Temperature effect on sensitivity (%/°C): 	≤ ± 0.0166%

Excitation voltage: 	5 to 15 Volts

Zero balance (%RO): 	≤ ± 5.0%

Insulation resistance (megohms): 	≥5000

Compensated temperature range (°C):  	–10o to +40o

Operating temperature range (°C): 	–40o to +80o

Safe load limit (% of Capacity): 	200%

Ultimate load (% of Capacity): 	300%

Load cell material: 	Stainless steel

Sealing: 	Hermetically sealed; cable entry sealed by glass to metal header

Protection: 	NEMA 6P/IP69K

The scale shall have self diagnostic capabilities able to identify load cell problems, failure, and predict failure before it occurs to prevent downtime. The diagnostics are to measure load cell counts (not weight) and will be used to determine reliability. Should a load cell fail, the instrumentation shall identify the specific load cell that has failed. All trouble shooting shall be done from within the scale house.

Load cell output shall be a DC analog signal. Load cells with digital output shall not be acceptable. Each load cell shall be optically isolated via a smart sectional controller. PC boards shall be encapsulated in epoxy or similar material. A board that is not protected in this fashion is unacceptable. Furthermore, each encapsulated board shall be housed in a type 304 Stainless steel enclosure rated NEMA 4X. Access to the encapsulated board within the smart sectional controller enclosure shall be achieved without the use of tools.  Bolts, screws or other hardware shall not be used to seal the smart sectional controller enclosure.

  - Platform and Foundation Requirements. The weighbridge and load cell assemblies shall be supported by a reinforced concrete pier type or full slab foundation as indicated on the drawings. The dimensions for the scale foundation and platform shall be as recommended by the equipment manufacturer and accepted by the Engineer.  Reinforcing steel placement and structural steel embedment placement shall be performed as shown on the manufacturer's foundation drawings.

Maximum distance from the top of the base slab to the base of rails shall not exceed 21 inches.

The scale manufacturer shall furnish the following items for construction of the scale platforms and pits:

- Weighbridge modules

- Longitudinal and lateral bumper checking devices

- Load cells and load cell assemblies.

- Anchor bolts.

- Platform and endwall painted structural steel embedments.

- PART 3 – EXECUTION

  - Installation.  The scale shall be manufactured, provided, and installed by a scale company that has a minimum of five years of experience installing similar truck scale systems.

The installer shall configure the scale system as indicated on the certified drawings.  All concrete work shall be as specified in the cast-in-place concrete section.  Anchor bolts shall be set as required by the scale manufacturer’s drawings.

  - Manufacturer’s Field Services.  Where scheduled in the equipment schedule section, an experienced, competent, and authorized representative of the manufacturer shall provide field services for equipment furnished under this section.  Field services shall meet the requirements of Manufacturer's Field Services in the quality control section.

  - Field Testing and Acceptance.  An authorized manufacturer’s representative shall provide the required scale certification for capacity and accuracy to the Engineer as required by the applicable state department of weights and measures and any other applicable state or county agency.

  - Personnel Training.  An experienced, competent, and authorized representative of the manufacturer shall train the Owner's personnel in operating, maintaining, and repairing the equipment specified in this section.  The training provided shall meet the requirements of Personnel Training Services in the quality control section.  The number of training sessions and duration of each session shall be as scheduled in the equipment schedule section.

End of Section